GAS CYLINDER MONITOR V1

Comprehensive Session Record

2026-06-18  |  Full Day Session

Gratian Technologies  |  Project13

This document covers everything accomplished in the 2026-06-18 session: hardware diagnosis, all code fixes applied, architecture decisions locked, boot sequence redesign, cal_factor design for V1, observability design, WebUI plans, and complete node backlog status.

# 1. Where We Started This Session

At session start, Node Layer 1 was complete (1A+1B+1C+1D) and verified on hardware. The hub skeleton was deployed at arduino@AQ3:7000. The WebUI was showing live weight via BLE. No gas logic, no log persistence, no boot sequence redesign yet.

The WebUI showed -49g weight and -2% remaining. Three separate bugs were visible:

- Platform was not empty during boot — tare zeroed a load, so everything lighter read negative

- TODO 1B-stuck: tare_variance_raw always 0.0f — stuck check fired on every boot → DEGRADED

- Gas% formula used current/starting instead of (gross - steel) / 14200 — downstream of bug 1

The hardware diagnostic sketch (hw_test) was run. Results from two runs:

| Test | Run 1 Result | Run 2 Result |
| --- | --- | --- |
| Raw stability | PASS — STD=100, CV=0.094% | PASS — STD=140, CV=0.132% |
| Cell 1 lift | PASS — delta=-95.7g | WARN — delta=+25.1g (polarity) |
| Cell 2 lift | WARN — delta=-34.6g | WARN — delta=+60.4g (polarity) |
| Cell 3 lift | WARN — delta=-14.7g | WARN — delta=+41.6g (polarity) |
| Load test 1000g | PASS — 1004.3g, cal=36.25 | PASS — 1004.8g, cal=36.27 |

Conclusion: Hardware is healthy. Lift test contradictions between runs were caused by platform geometry and mechanical redistribution — not cell faults. Load test (Section 4) is the authoritative test. Both runs agreed: 1004g for 1000g placed. Hardware confirmed healthy.

# 2. All Bugs Found and Fixed This Session

## 2.1 Bug: sigma=25.62g causing 48% weight reading

Root cause: one load cell had an intermittent connection overnight. The NOISE phase ran with one cell dropping in and out between samples — producing enormous variance. Sigma=25.62g vs healthy 4-6g. The fix was physical: reconnect the loose wire at the junction. Hardware verified clean on next boot (sigma=4.74g, 1000g reads 997-1002g).

## 2.2 Bug: NOISE result=WARN on every healthy boot (F1)

Root cause: NOISE sigma thresholds were calibrated for the old single-cell STM32 platform (healthy sigma ~1.87g). On 3-cell ESP32-C3, healthy sigma is 4.68-5.44g. The old PASS gate was below the healthy range, so it fired WARN every boot even on perfect hardware.

Fix applied to noise.cpp:

static const float NOISE_SIGMA_PASS_G = 8.0f;   // was ~3.0f

static const float NOISE_SIGMA_WARN_G = 15.0f;  // was ~6.0f

Rationale: healthy max is 5.33g. PASS gate at 8.0g = 1.5x margin. One open cell causes sigma ~25g. WARN gate at 15.0g catches degraded state. FAIL at >=15.0g indicates hardware fault.

## 2.3 Bug: quality=DEGRADED on every healthy boot (F2)

Root cause: health_check() stuck detection reads tare_variance_raw from TareResult struct. tare.cpp never populates this field — always 0.0f. Stuck check compares 0.0f against threshold → always 'stuck' → always DEGRADED. False positive from incomplete struct field.

Fix applied to health.cpp:

bool skip_stuck = (tare_variance_raw == 0.0f);

if (!skip_stuck) { /* existing stuck check */ }

TODO 1B-stuck remains: properly expose tare_variance_raw from tare.cpp through TareResult struct. Fix is deferred — the guard prevents false firing until the proper fix is built.

## 2.4 Bug: NOISE result=WARN persisted after threshold fix

Root cause: the orchestrator in gas_monitor_v1.ino checks noise result using strstr(nr.diagnosis, 'too high') pattern. But the NOISE phase ran with g_cal_factor=0.0f (not yet derived), so noise_update() took the raw-counts branch: grams = raw - tare_raw (no division). Sigma came out in raw counts (~180 counts) which when compared against the 15.0g gram-based threshold always exceeded it — always 'too high' — always WARN.

Fix applied to gas_monitor_v1.ino — load saved cal_factor from SPIFFS before NOISE phase:

case STATE_NOISE: {

if (g_cal_factor == 0.0f) {

float _sc = cal_load_last();

if (_sc > 0.0f) g_cal_factor = _sc;

}

NoiseResult nr = noise_update(...);

## 2.5 Bug: sigma=0.09g causing 200+ false WEIGHT_EVENTs (boot=33)

Root cause: noise_update() stored samples in grams (divided by cal_factor during accumulation). noise_recompute_sigma() then divided by cal_factor again after CAL. Double division caused sigma to shrink from ~5g to ~0.09g. Threshold = 4 × 0.09 = 0.36g. Every vibration, finger touch, and breath crossed this threshold.

Fix applied to noise.cpp — store raw counts in s_samples[], divide by cal_factor exactly once at sigma computation:

// BEFORE (wrong):

grams = ((float)raw - tare_raw) / cal_factor;

s_samples[s_count] = grams;

// AFTER (correct):

float raw_counts = (float)raw - tare_raw;

s_samples[s_count] = raw_counts;  // raw only

// ...

float sigma_g = (cal_factor > 0.0f) ? (sigma_raw / cal_factor) : sigma_raw;

## 2.6 Bug: false WEIGHT_EVENT REMOVED with nothing touched (boot=31, line #0045)

Root cause: delay-line window was 20 ticks (2 seconds). Two consecutive 2-second rolling means could differ by >21.76g (4x sigma) purely from natural load cell creep on a static 1000g load. The platform had not fully settled mechanically after 18+ minutes.

Fix applied to weight.cpp:

#define BUF_SIZE 40  // was 20

4-second comparison window. SE of 40-sample mean = sigma/sqrt(40) = 5.44/6.32 = 0.86g. Natural mean difference drops from ~10g worst case to ~7g. Threshold at 21.76g — no false triggers.

## 2.7 Bug: NimBLE onWrite signature mismatch

Root cause: NimBLE-Arduino by h2zero changed the callback signature between versions. The generated code used the old single-parameter form. esp32 v3.0.7 ships NimBLE v1.4+ which requires two parameters.

// Old (wrong for NimBLE v1.4+):

void onWrite(NimBLECharacteristic* c) override

// Correct for NimBLE v1.4+:

void onWrite(NimBLECharacteristic* c, NimBLEConnInfo& connInfo) override

Fixed via sed on AQ3 before SCP to laptop.

# 3. Hardware-Verified Values This Session

| Parameter | Value | Status |
| --- | --- | --- |
| cal_factor (boot=29) | 36.2689 raw/g (loaded SPIFFS) | VERIFIED |
| cal_factor (boot=35 hw_test) | 36.25-36.27 raw/g derived | VERIFIED |
| sigma (boot=31) | 5.44g | VERIFIED — healthy |
| sigma (boot=35) | 3.16g | VERIFIED — healthy |
| sigma (boot=33 — bug) | 0.09g | VERIFIED — double-division bug |
| threshold (boot=31) | 4 × 5.44 = 21.76g | DERIVED |
| threshold (boot=35) | 4 × 3.16 = 12.64g | DERIVED |
| BUF_SIZE (delay-line) | 40 ticks = 4 seconds | LOCKED |
| 1000g accuracy (boot=31) | 994-1003g over 30 heartbeats | VERIFIED ±6g |
| 1000g accuracy (boot=35) | 986.5g first HB | VERIFIED |
| SETTLE duration | ~2.1s | VERIFIED |
| TARE duration | ~21.1s (N=200 at 10 SPS) | VERIFIED |
| NOISE duration | ~20.1s (N=200 at 10 SPS) | VERIFIED |
| TARE_WAIT timeout | 60s | LOCKED |
| Boot time with TARE_WAIT | ~103.9s total | VERIFIED |
| tare_raw boot=31 | -105687.7 | VERIFIED — saved to SPIFFS |
| tare_raw boot=33 | -105574.4 | VERIFIED |
| tare_raw boot=35 | -105232.4 | VERIFIED |

# 4. Features Built This Session

## 4.1 BLE Command Characteristic

A writable BLE characteristic added to the node allowing hub-to-node commands. All hub commands flow through one characteristic using ASCII string format.

| Command | Purpose | Status |
| --- | --- | --- |
| TARE | Hub commands fresh tare execution | Built |
| SKIP_TARE | Hub commands load saved tare from SPIFFS | Built |
| SET_CAL:<value> | Hub sends cal_factor, node skips interactive CAL | Built |
| RETARE | Hub triggers fresh retare during RUNNING state | Built (stub) |
| DUMP_LOG | Hub requests SPIFFS journal file transfer | Stub only |
| CLEAR_LOG | Hub confirms file received, node clears SPIFFS | Stub only |

UUIDs locked:

Command char: c8a2f1e3-4d6b-4a7c-8e9f-1b2d3e4f5a6b (write-without-response)

Log char:     d7b3e2f4-5e7c-4b8d-9f1a-2c3e4f5a6b7c (notify — reserved for 1E)

## 4.2 STATE_TARE_WAIT — New Boot State

Inserted between SETTLE and TARE in the boot sequence. Node waits up to 60 seconds for hub to send TARE or SKIP_TARE command. 60-second timeout is fail-safe — triggers fresh tare and continues. Hub is the decision maker; node never decides tare path alone.

Boot sequence (new):

SETTLE → TARE_WAIT → TARE → NOISE → CAL → RUNNING

Verified on boot=31 and boot=33: TARE_WAIT fires, 60s timeout runs, fresh tare executes, tare_raw saved to SPIFFS.

## 4.3 Tare SPIFFS Persistence

Two new functions added to tare.cpp:

tare_save_to_spiffs(float tare_raw) — saves to config.json after every fresh tare

tare_load_from_spiffs(float* out) — loads with sanity check [-200000, -50000]

Tare is now saved after every fresh derivation. SKIP_TARE loads this value. Cross-boot tare consistency is possible when hub sends SKIP_TARE.

## 4.4 STATE_RETARE

New state in RUNNING. Triggered by g_cmd_retare_pending flag set by RETARE command. Pauses weight readings, runs fresh N=200 tare, saves to SPIFFS, returns to RUNNING. Used by HUB-001 auto-retare on cylinder removal.

## 4.5 WebUI Auto-Launch Script (Windows)

A Python script (gas_monitor_webui_launcher.py) runs silently at Windows startup. Polls http://192.168.88.20:7000 every 5 seconds. When hub WebUI responds, opens the default browser automatically. Exits after opening — does not stay resident. Installed via Windows Startup folder shortcut using pythonw.exe (no console window).

# 5. Boot Sequence — Final V1 Design

## 5.1 Complete Boot Sequence

| Phase | Duration | Who decides | What happens |
| --- | --- | --- | --- |
| SETTLE | ~2.1s | Node autonomous | Platform stabilises. No load cells read. |
| TARE_WAIT | 0–60s | Hub decides | Node waits. Hub sends TARE or SKIP_TARE. |
| TARE | ~21s | Node executes | N=200 samples → mean. Saves tare_raw to SPIFFS. |
| NOISE | ~20s | Node autonomous | N=200 samples → sigma, threshold. No cal_factor needed. |
| CAL | ~0.1s | Hub via SET_CAL | Hub sends saved cal_factor. Node loads and saves to SPIFFS. |
| RUNNING | Continuous | Hub monitors | Weight readings via BLE notify. Journal via Serial + SPIFFS. |

## 5.2 Tare Logic — Problems and Mitigations

Key principle: tare is a snapshot in time. The real world is continuous. Load cells drift with temperature, platforms move, temperatures change.

| Problem | Severity | Status | Mitigation |
| --- | --- | --- | --- |
| Temperature zero drift | Medium | V2 | ~6g/hr drift. Periodic retare during known-empty moments. |
| Platform mechanical disturbance | High | V2 | HUB-002 disturbance detection (not yet built). |
| Cylinder placed during TARE_WAIT | Medium | Partial | Hub confirms empty before sending TARE command. |
| Stale SPIFFS tare on SKIP_TARE | Low-Med | Planned | Hub validates gross after SKIP_TARE command sent. |
| Retare with new cylinder already on | High | Planned | Hub adds settle confirmation after RETARE. |
| Hub off, weight on platform at boot | Medium | Planned | Node post-tare self-check (next build — N-TARE-CHECK). |

# 6. Cal Factor — Complete V1 Design

## 6.1 What Cal Factor Is

Cal factor is the conversion constant between raw ADC counts and grams: grams = (raw - tare_raw) / cal_factor. It encodes everything about the physical system: cell sensitivity, HX711 gain, wiring resistance, mechanical load distribution. On this 3-cell platform, verified range is 34.47–36.27 raw/g. It drifts ±3% with kitchen temperature range.

Critical rule: tare and cal_factor are always a PAIR. They must come from the same physical session. Hub is responsible for keeping this pair together in config.json.

## 6.2 The Four Production Cases

### Case 1 — Brand new everything, first ever boot

- Node: freshly flashed, SPIFFS empty

- Hub: freshly deployed, no config.json

- Platform: empty, cylinder not yet placed

- Flow: SETTLE → TARE_WAIT → hub sends TARE → TARE → NOISE → CAL_WAIT

- User places fresh cylinder (known brand)

- Hub detects gross >26kg → looks up steel from company table

- Hub computes: cal_factor = (gross_raw - tare_raw) / (steel_kg + 14200g)

- Hub saves cal_factor + tare_raw as PAIR to config.json

- Hub sends SET_CAL to node → node saves to SPIFFS → RUNNING

- Accuracy: excellent — 29.5kg reference weight (130x better than 227g weight)

### Case 2a — Power cycle, platform empty

- Hub boots → checks SQLite → sends TARE (platform appears empty)

- Node tares → saves new tare_raw to SPIFFS

- Hub has cal_factor in config.json → sends SET_CAL

- Hub cross-checks first gross reading within ±5% of expected → accept silently

### Case 2b — Power cut, cylinder already on platform

- Hub boots → knows cylinder was TRACKING → sends SKIP_TARE

- Node loads tare_raw from SPIFFS (previous session value)

- Hub sends SET_CAL with SAME cal_factor paired with that tare

- Critical: same-session values must travel together — hub keeps them as a pair

- Hub cross-checks gross against saved steel + remaining gas estimate (±10% tolerance)

### Case 3 — Reboot, platform empty, hub already running

- Identical to Case 2a. Hub connects faster, sends TARE before timeout.

### Case 4 — Cylinder removed, fresh cylinder placed

- WEIGHT_EVENT REMOVED → gross drops to ~0 → hub state = UNINSTALLED

- Hub monitors settle gate → sends RETARE → node retares → saves new tare_raw

- User places new cylinder → hub detects gross >26kg

- Hub derives NEW cal_factor from new cylinder using lookup table

- Cross-check against old cal_factor: within 5% → update, outside 5% → WARN WebUI

- Accept new value regardless — cylinder IS the ground truth

## 6.3 The Governing Rules

Three rules emerge from the four cases:

| Rule | Detail |
| --- | --- |
| Tare + cal_factor are a PAIR | Hub stores both together with timestamp. If tare is loaded from SPIFFS, the matching cal_factor must be sent — not a freshly derived one. |
| Cylinder IS the calibration reference | Fresh cylinder at install = 29.5kg reference. 130x better than 227g development weight. System recalibrates automatically on every cylinder replacement. |
| Hub derives, node stores and uses | Node never derives cal_factor from a known weight in production. Hub computes from BIS anchor. Node receives via SET_CAL, saves to SPIFFS, uses for conversion. |

## 6.4 Edge Case — First Boot, Hub Off

Node boots, no hub, no SPIFFS value. Currently node could wait indefinitely for interactive CAL input. Fix needed: 120s timeout → use hardcoded fallback 36.0 raw/g (known good for this platform) → flag FAULT in journal → enter RUNNING with DEGRADED quality. Hub connects later → sends SET_CAL with real value → node updates immediately.

## 6.5 Cal Factor Problems and Mitigations

| Problem | Severity | Mitigation |
| --- | --- | --- |
| Stale SPIFFS cal_factor (temp drift) | Medium | Accepted V1. Max 5% error. Delta measurements immune (drift cancels in subtraction). |
| Small development weight amplifies error | Medium | Accepted V1. Production uses 29.5kg cylinder reference — 130x more accurate. |
| No cross-check on loaded cal_factor | High | Planned hub-side: compare gross reading against expected from lookup table on every boot. |
| Hub off, no SPIFFS, node hangs in CAL | High | 120s timeout fallback to 36.0 raw/g — next node fix. |
| Tare and cal_factor from different sessions | Medium | Hub sends paired values together. SKIP_TARE always accompanied by matching SET_CAL. |

# 7. Observability and Log Architecture — Locked Design

## 7.1 The Problem

In production, no Serial Monitor is attached. Journal lines emit to Serial only — they vanish. Post-mortem diagnosis of node failures is impossible without log persistence and retrieval. The hub also has no structured event log. Flying blind in production.

## 7.2 Node Log Transport — Two Triggers

| Trigger | What happens |
| --- | --- |
| Size threshold (node-initiated) | When SPIFFS file reaches 25KB, node sets transfer_pending flag. On next RUNNING tick, streams file to hub via BLE log characteristic. No hub request needed. |
| Hub request (hub-initiated) | Hub detects new boot number on BLE connect. Hub sends DUMP_LOG command via write characteristic. Node transfers immediately regardless of file size. |

## 7.3 Transfer Protocol

- Node notifies LOG_START sentinel on log characteristic

- Node reads /node_journal.log line by line, notifies each line as one BLE packet

- Node notifies LOG_END sentinel

- Hub receives LOG_END → renames temp file to logs/node/node_YYYY-MM-DD_bootNN.log

- Hub sends CLEAR_LOG command ONLY after file confirmed written to disk

- Node receives CLEAR_LOG → deletes /node_journal.log → resets RAM size counter

- If BLE drops mid-transfer: node keeps file. Hub discards .tmp. Next connect → hub requests again.

Transfer blocked during: SETTLE, TARE_WAIT, TARE, NOISE, CAL phases. Only allowed in STATE_RUNNING. Size threshold hit during boot → set pending flag → transfer when RUNNING begins.

File size tracking: RAM counter g_journal_file_bytes incremented by strlen(line) on every append. On boot, initialised from actual SPIFFS file size (handles reboot without clear).

## 7.4 Log Directory Structure on AQ3

~/ArduinoApps/gas-cylinder-monitor/

└── logs/

├── node/

│   ├── node_2026-06-18_boot031.log  ← one per boot session

│   ├── node_2026-06-18_boot033.log

│   └── .transfer.tmp               ← deleted on complete

└── hub/

├── hub_2026-06-18.log           ← one per day, appended

└── hub_2026-06-17.log

## 7.5 WebUI Serial Feed Design — Three Tabs

All journal events visible on WebUI without Serial Monitor. Three tabs:

- Live tab: current weight, quality badge, gas%, boot sequence phase summary with durations and results

- Boot log tab: full boot sequence for current session — colour-coded BOOT/CMD tags, all phases with timing

- Serial feed tab: real-time stream of all journal lines as they arrive from node — monospace font, colour tags [BOOT]=blue, [RUN]=green, [HB]=grey, [FAULT]=red, [CMD]=purple

- Command input box in Serial feed tab: type TARE / RETARE / DUMP_LOG and send via BLE command char

Cal factor display on WebUI: Path taken (A=loaded/B=derived), value, cross-check delta (green <5%, amber >5%), last update boot number.

# 8. Complete Node Backlog — All Remaining Work

## 8.1 Code Fixes

| ID | What | Fix required | Priority |
| --- | --- | --- | --- |
| TODO 1B-stuck | tare_variance_raw always 0.0f — stuck check never meaningful | Expose variance from tare.cpp through TareResult struct | LOW |
| TODO 1B-persistence | prev_cal_factor / prev_sigma_g not read from config.json at boot | Read/write at boot and after CAL_SUCCESS | MEDIUM |
| CAL timeout | No timeout on interactive CAL — node hangs if hub off and SPIFFS empty | 120s timeout → 36.0 fallback → DEGRADED quality | HIGH |

## 8.2 Features to Build (in dependency order)

| ID | What | Why | Gate |
| --- | --- | --- | --- |
| N-TARE-CHECK | Post-tare self-check: detect weight on platform | Node detects if tare was taken with load — uses SPIFFS tare as fallback | No gate — next build |
| N1 | Journal → SPIFFS file + RAM size counter | Log survives without Serial in production | No gate — next build |
| 1E | BLE log notify char — streams journal lines live | Real-time observability without Serial | After BLE-CMD |
| N-LOG-TRANSFER | SPIFFS file transfer FSM — DUMP_LOG/CLEAR_LOG/25KB auto-push | Hub pulls full session log | After 1E |
| 1F | TARE_WAIT built — hub sends TARE/SKIP_TARE | Power cut safety (already built) | COMPLETE |
| N-TARE-SPIFFS | Save tare to SPIFFS after every fresh tare | Required for SKIP_TARE (already built) | COMPLETE |
| N-SET-CAL | SET_CAL command handler | Hub sends cal_factor (already built) | COMPLETE |
| HUB-001 node | RETARE command handler | Fresh tare on cylinder removal | COMPLETE (stub) |

## 8.3 Experiments Remaining

| ID | What | What it answers | Gate |
| --- | --- | --- | --- |
| 3E-008 | Temperature drift | How much does cal_factor and zero shift with kitchen heat? | After N1+1E |
| 3E-009 | Long-run 6hr stability | Characterise 190g drift. Gas% accuracy budget over duty cycle. | After 3E-008 |
| 3E-010 | Load cell failure injection | Does health module detect and flag open cell correctly? | After 3E-009 |
| 3E-005 | V1 anchor validation | End-to-end proof: hub derives steel, gas% goes live. | After hub Group 4 |

# 9. Build Order — Next Sessions

| Order | Item | Detail |
| --- | --- | --- |
| 1 — NOW | N-TARE-CHECK | Post-tare self-check. Node detects weight on platform at boot. Uses SPIFFS saved tare as fallback. |
| 2 | N1 — Journal → SPIFFS | Every journal line appended to /node_journal.log. RAM counter tracks size. Transfer pending at 25KB. |
| 3 | BLE-CMD log char | 1E log notify characteristic starts streaming lines. Hub receives live journal feed. |
| 4 | N-LOG-TRANSFER | DUMP_LOG/CLEAR_LOG full FSM. LOG_START / LOG_END sentinels. Hub saves to logs/node/. |
| 5 | Hub log directory | Hub-side event log. logs/hub/hub_YYYY-MM-DD.log. BLE events, readings, state changes. |
| 6 | WebUI serial feed panel | Three tabs. Boot log, live feed, command input. Real-time journal on WebUI. |
| 7 | CAL timeout fix | 120s timeout on interactive CAL wait → 36.0 fallback → DEGRADED. |
| 8 | 3E-008 temperature drift | Kitchen heat cycle. Measure cal_factor and zero drift. |
| 9 | 3E-009 long-run stability | 6hr static load. Characterise drift budget. |
| 10 | 3E-010 failure injection | Disconnect one cell. Verify health module catches it. |
| 11 | Hub Group 4 domain logic | Company lookup table, BIS anchor, cal_factor derivation, state machine UNINSTALLED/TRACKING/LOW_GAS. |
| 12 | 3E-005 anchor validation | End-to-end: water container, hub derives steel, gas% live. |
| 13 | Hub Group 5 analytics | Burn rate from real data only. Days remaining. |
| 14 | WebUI full dashboard | Gas gauge, trend, prediction, boot visibility panel. |
| 15 | Overnight integration test | 8hr full system test. All components. |

# 10. Architecture Decisions Locked This Session

| Decision | What was locked |
| --- | --- |
| V1 cal_factor source | Cylinder itself via BIS 14.2kg anchor. Hub lookup table provides steel weight. No small test weights in production. |
| Hub owns cal derivation | Node never derives cal_factor in production. Hub computes from gross reading + lookup table. Sends via SET_CAL. |
| Tare + cal_factor as PAIR | Always stored together with timestamp in hub config.json. SKIP_TARE always accompanied by matching SET_CAL from same original session. |
| Log transport — two triggers | Node auto-pushes at 25KB. Hub requests on new boot number detected. Same transfer mechanism for both. |
| Transfer only in RUNNING | Log transfer blocked during all boot phases. Timing-sensitive phases must not be interrupted by BLE streaming. |
| Hub clears node only after disk confirm | CLEAR_LOG sent only after hub confirms file written to disk. Node never loses log data on BLE drop. |
| WebUI three-tab design | Live (weight + boot phases), Boot log (full session), Serial feed (live journal stream + command input). |
| Production WebUI on HDMI screen | Development: laptop browser auto-opens. Production: Chromium kiosk mode on AQ3 HDMI output (same pattern as youtube-display project). |
| User picks company at install | One-time setup. User selects brand logo. Hub looks up steel weight from company table. No other user input required. |
| V2 scope — unknown partial cylinder | Cylinder replaced with unknown brand and unknown remaining gas is V2/V3. V1 assumes fresh cylinder at replacement. |
| Drift and disturbance — V2 | Temperature zero drift and platform mechanical disturbance detection (HUB-002) are V2. V1 accepts these risks within the characterised budget. |
| BUF_SIZE = 40 ticks | 4-second delay-line comparison window. Reduces false WEIGHT_EVENTs from load cell creep on static loads. |

# 11. Key Learnings This Session

L-056: NOISE thresholds must be re-evaluated when platform changes. Single-cell STM32 sigma ~1.87g vs 3-cell ESP32-C3 sigma 4.68-5.44g. Always re-derive gates from hardware characterisation, never carry forward from different platform.

L-057: When loading a saved cal_factor before NOISE phase, samples must still be stored as raw counts. Cal_factor must be applied exactly once — at sigma_raw → sigma_g conversion. Storing grams causes double-division when noise_recompute_sigma() runs after CAL.

L-058: NimBLE-Arduino v1.4+ changed onWrite callback signature to include NimBLEConnInfo& parameter. Always check library changelog when compile error says 'does not override' on a callback method.

L-059: Lift test on 3-cell parallel platform is unreliable as a cell health test. Corner lift redistributes load to other cells, and net effect on total weight depends on platform geometry. Section 4 load test (known weight, measured output) is the authoritative health check.

L-060: Cal_factor and tare are a PAIR derived in the same physical state. Loading a tare from SPIFFS (one temperature/session) while using a freshly derived cal_factor (different temperature/session) introduces systematic error. Hub must keep and send both values together.

L-061: The cylinder itself is the best calibration weight. At 29.5kg, measurement error contribution to cal_factor is 130x smaller than using a 227g development weight. Production calibration via BIS 14.2kg anchor is inherently more accurate than any development calibration.

L-062: Gas consumption tracking (delta measurements) is immune to cal_factor drift because drift affects the absolute reading equally at the anchor point and the current reading — it cancels in subtraction. Only absolute gas weight is affected by drift.

L-063: Delay-line window of 20 ticks (2 seconds) is insufficient for platforms with ongoing mechanical creep. At sigma=5.44g, two independent 2-second means can differ by >threshold purely from natural variance on a static load. 40-tick window (4 seconds) reduces false trigger rate to near-zero.

# 12. BLE UUIDs — All Locked

| Item | UUID / Value | Status |
| --- | --- | --- |
| Service UUID | aa206b91-235b-42aa-b370-453a3feedf35 | LOCKED |
| Weight char (notify) | b9b25bb1-f2a9-4545-b48f-295ab2789f41 | LOCKED |
| Command char (write-NR) | c8a2f1e3-4d6b-4a7c-8e9f-1b2d3e4f5a6b | LOCKED — built this session |
| Log char (notify) | d7b3e2f4-5e7c-4b8d-9f1a-2c3e4f5a6b7c | LOCKED — registered, streaming not yet built |
| Device name | GasCylMonitor | LOCKED |
| Node MAC | 10:00:3B:CD:63:32 | VERIFIED |
| Payload format | {"grams":29420.5,"quality":"GOOD","sigma":2.3} | LOCKED |
| Command format | ASCII string e.g. TARE\n SET_CAL:36.2689\n | LOCKED |

# 13. Node Status — End of Session

| SETTLE | PASS — 2.1s consistent |
| --- | --- |
| TARE_WAIT (1F) | PASS — 60s timeout working |
| TARE + SPIFFS save | PASS — N=200, saves tare_raw |
| NOISE result | PASS — OK on boot=35 (sigma=3.16g) |
| CAL (SET_CAL path) | PASS — loads from SPIFFS, hub path built |
| quality=GOOD | PASS — healthy hardware shows GOOD |
| False WEIGHT_EVENTs | FIXED — BUF_SIZE=40, zero false triggers |
| Journal format | PASS — all events correct format |
| BLE command char | PASS — compiled, running |
| Journal → SPIFFS (N1) | NOT BUILT — next session |
| 1E log streaming | NOT BUILT — after N1 |
| CAL timeout fallback | NOT BUILT — hub-off deadlock risk |
| TODO 1B-stuck | GUARDED — false positive prevented |
| TODO 1B-persistence | DEFERRED — low priority |

End of Session Record 2026-06-18

Gratian Technologies / Project13 — Gas Cylinder Monitor V1