Gas Cylinder Monitor V1

Load Cell Platform — Operations & Tare Mechanism Reference

Gratian Technologies · Project 13 · Gas Cylinder Monitor

Purpose: everything you need to know about the physical platform — how tare works, what happens when you add or remove things, how to swap the platform, and what is safe vs dangerous.

# 1. The Core Principle — How tare_raw Works

Understanding this one principle makes all the rules below obvious. The tare is a measurement of the platform at a moment in time with nothing on it. Everything on the platform at tare time becomes permanently invisible to the system.

The formula the system uses for every reading is:

gross_g = (raw_reading - tare_raw) / cal_factor

# tare_raw was measured when platform was empty.

# Everything on the platform at tare time = zero forever.

# Gas consumption = gross_g(now) - steel_g

# steel_g was derived at anchor = gross_g when cylinder was full

This means any permanent fixture on the platform at tare time is absorbed into tare_raw and becomes mathematically invisible. It never shows up in gross_g. It never affects gas% or consumption calculations. This is not a bug — it is by design.

THE GOLDEN RULE

Whatever is on the platform when tare is taken defines 'zero'.

Everything permanent on the platform at tare time = invisible forever.

Everything added after tare = shows up in readings as real weight.

Everything removed after tare = shows up as negative weight error.

# 2. What Is Safe to Keep on the Platform

The following items are safe to keep on the platform permanently because they are present at tare time and therefore absorbed:

## 2.1  Always safe — present at tare, permanent

| Item | Why safe | Effect on readings |
| --- | --- | --- |
| Rubber mat on top plate | Present when tare is taken. Part of platform assembly. | Zero — absorbed into tare_raw |
| Drip tray (if placed before tare) | Present when tare is taken. | Zero — absorbed into tare_raw |
| Platform plate itself | Always present at tare. | Zero — absorbed into tare_raw |
| Any protective sheet placed before tare | Absorbed into tare_raw at tare time. | Zero — absorbed into tare_raw |
| Wires and junction block under platform | Part of hardware assembly, present at tare. | Zero — absorbed into tare_raw |

## 2.2  Dangerous — added or removed after tare

| Action | What happens | Severity |
| --- | --- | --- |
| Add a mat/sheet/tray AFTER tare | Its weight shows as permanent positive offset in gross_g. System thinks cylinder is heavier than it is. gas% reads high. | HIGH |
| Remove a mat/sheet/tray AFTER tare | Its weight shows as permanent negative offset in gross_g. System thinks cylinder is lighter than it is. gas% reads low. | HIGH |
| Place cylinder then take tare | Cylinder weight absorbed into tare_raw. System reads 0g gross with cylinder on platform. Everything breaks. | CRITICAL |
| Move platform slightly after tare | Mechanical shift changes cell loading geometry. Introduces unknown zero offset. Silent wrong readings. | HIGH |
| Rest something on platform edge | Extra weight reads as part of gross_g. System thinks more gas remains than there is. | MEDIUM |

⚠ CRITICAL: Adding or removing anything from the platform after tare requires a fresh tare.

There is no software fix for a physical change made after tare.

The only correct action is: remove cylinder → remove/add the item → take fresh tare → reinstall cylinder.

# 3. Adding a Sheet or Mat on the Platform

Sometimes you want to add a protective sheet, anti-vibration pad, or drip tray on the platform. Here is exactly how to do it correctly depending on whether the system is already deployed.

## 3.1  Before first deployment (easiest case)

If the system has never been deployed and tare has never been taken, simply place the sheet on the platform before the first tare. It will be absorbed and invisible forever.

- Place sheet/mat on platform top

- Power on node with platform EMPTY (nothing else on it)

- Hub sends TARE command — node takes fresh tare with sheet on platform

- Sheet is now part of zero forever

- Place cylinder normally

## 3.2  After deployment (system already running)

⚠ Do NOT simply place a sheet on the platform while the cylinder is there.

The sheet weight will appear as gas_g increase. The system will think gas was added. State machine may wrongly trigger a refill detection.

The correct procedure when adding a sheet to a live deployed system:

- Step 1: Remove the gas cylinder from the platform

- Step 2: Wait for REMOVAL_GRACE_S (120 seconds) to pass — system transitions to UNINSTALLED state

- Step 3: Wait an additional 60–90 seconds for viscoelastic creep recovery (cells need to spring back)

- Step 4: Place the sheet/mat on the empty platform

- Step 5: In the WebUI, click 'Install cylinder' — this triggers fresh tare in the hub flow

- Step 6: The hub sends TARE command to node — fresh tare captures sheet as zero

- Step 7: Place the cylinder on the platform

- Step 8: Anchor collection begins (BOOTSTRAP_ANCHOR state) — gas% goes live after 5 stable readings

Why the 60–90 second wait after removal matters: the load cell beam deflects under sustained load (viscoelastic creep). After removal, it springs back but overshoots slightly, then settles. Taking tare before it settles introduces a small negative offset. Wait for full mechanical recovery before tare.

# 4. Swapping or Replacing the Platform

Platform swap means replacing the physical assembly — the plates, load cells, or both. This is a significant operation because cal_factor and tare_raw are both derived from the physical assembly. A new platform = new physics = must re-derive both.

## 4.1  What changes and what does not

| Parameter | Changes on platform swap? | Reason |
| --- | --- | --- |
| tare_raw | YES — must re-derive | New platform has different mechanical zero. Old tare_raw is invalid. |
| cal_factor | YES — must re-derive | New cells have different sensitivity. Old cal_factor is invalid. Even same model cells vary ±3%. |
| NET_GAS_G | No — unchanged | BIS IS 3196 legal constant. 14.2kg for all Indian domestic cylinders. |
| All domain constants | No — unchanged | State machine logic is independent of physical platform. |
| steel_g | YES — re-derived automatically | New tare_raw + new cal_factor = new gross_g reading. Anchor must fire again to re-derive steel_g. |

## 4.2  Platform swap procedure

⚠ Never reuse tare_raw or cal_factor from the old platform on a new platform.

They are physically coupled to that specific assembly. Using old values on new hardware gives wrong readings with no error indication.

Complete procedure for platform swap:

- Step 1: Remove the gas cylinder. Wait for UNINSTALLED state (120s grace window).

- Step 2: Power off the node (unplug USB).

- Step 3: Swap the physical platform assembly.

- Step 4: Level the new platform with a spirit level. Target: within 0.5° of horizontal. 1° tilt = up to 297g cross-axis error on a full cylinder.

- Step 5: Power on the node. Hub will send TARE command (cylinder_state = UNINSTALLED). Node takes fresh tare on new platform.

- Step 6: Hub must send SET_CAL with a known-good cal_factor for the new platform, OR the node uses 36.0 fallback (DEGRADED) until re-calibrated.

- Step 7: If new cells: re-derive cal_factor using a known weight (see calibration procedure). Send new cal_factor via SET_CAL.

- Step 8: Place cylinder. Click 'Install cylinder' in WebUI. Anchor fires — new steel_g derived automatically from new gross_g reading.

If same cell type (YZC-161A) but new assembly: expect cal_factor to be in the 34–38 raw/g range. Use 36.0 as starting fallback. Derive precisely with a known weight before production.

## 4.3  Levelling rules

- Always level the base plate on the kitchen floor before deployment.

- Use a spirit level on the top plate surface (not the base plate — the rubber feet may not be equal).

- Target: within 0.5° of horizontal. 1° tilt adds ~1–3% cross-axis error.

- Lock the adjustable rubber feet with locknuts once level is confirmed.

- Re-check level whenever platform is moved, even a few centimetres.

- After moving platform: always take fresh tare before placing cylinder.

# 5. Viscoelastic Creep — Why You Must Wait After Removal

This is one of the most important physical phenomena to understand. Load cell beams are metal. Metal is not perfectly elastic — under sustained load, atoms slowly rearrange (viscoelastic relaxation). This causes the reading to drift over time even when nothing changes.

## 5.1  Creep under load (cylinder on platform)

| Time after placement | What happens | Effect on reading |
| --- | --- | --- |
| 0–30 seconds | Elastic deflection. Beam springs to correct position. | Reading accurate. |
| 30s – 4 hours | Viscoelastic creep begins. Beam slowly deforms further. | Reading drifts down 30–80g from initial. Looks like gas consumed. |
| 4+ hours | Creep plateaus. Beam reaches new equilibrium. | Reading stable. This stable offset is constant for the rest of the cylinder life. |
| 30–45 days later (removal) | Beam springs back elastically, then slowly recovers over 10–20 minutes. | Reading briefly incorrect after removal. Must wait before fresh tare. |

Key insight: the creep offset is constant throughout the cylinder cycle. Gas consumption = gross_g(full) − gross_g(now). Since both readings have the same creep offset baked in, the offset cancels out. Creep does not affect consumption accuracy — it only affects the absolute gas_g reading by ~0.35%.

## 5.2  Creep recovery after removal — the waiting rule

NEVER take tare immediately after removing a cylinder that has been on the platform for days or weeks.

The beam is mid-recovery. A tare taken now will have a negative offset baked in.

Wait at least 60–90 seconds minimum. After a full 30-45 day cylinder: wait 10–20 minutes ideally.

- After removing a cylinder that was on the platform < 1 hour: wait 60 seconds before tare.

- After removing a cylinder that was on the platform 1–24 hours: wait 5 minutes before tare.

- After removing a cylinder that was on the platform > 24 hours (production case): wait 10–20 minutes before tare.

- The system's auto-tare stability window already handles this — it waits for inter-reading drift < 1.0g before accepting tare. Trust this mechanism in production.

# 6. Platform Disturbance — What Causes Silent Errors

A platform disturbance is any physical change to the platform geometry after tare has been taken. These are dangerous because the system has no way to detect them — it just reports wrong readings silently.

## 6.1  Common disturbance causes

| Disturbance | Typical zero shift | Detectable? |
| --- | --- | --- |
| Moving platform 1–2 cm on floor | 50–300g shift | No — silent |
| Someone stepping on platform edge momentarily | May cause permanent zero shift if cells stressed | No — silent |
| Cleaning under platform (nudging base) | 20–100g shift | No — silent |
| Gas regulator replacement (pressure on cylinder) | Temporary — recovers when released | No — silent during event |
| Platform feet loosening over time | Gradual drift over weeks | Detectable via HUB-002 (V2 feature) |
| Cable tension change (wire repositioned) | 5–30g | No — silent |

## 6.2  What to do after a suspected disturbance

- Remove the cylinder and wait for UNINSTALLED state (120 second grace window).

- Wait for creep recovery (60 seconds minimum).

- Click 'Install cylinder' in WebUI to trigger fresh tare.

- Place cylinder. Anchor fires and re-derives steel_g with new zero reference.

- System is now correct again.

⚠ If you suspect the platform moved with the cylinder still on it:

There is no way to correct this without removing the cylinder and re-anchoring.

The gas% shown may be wrong by an unknown amount until the cylinder is removed, re-tared, and re-installed.

# 7. tare_raw Drift Across Boots — Expected Behaviour

tare_raw is measured fresh on every boot where hub sends TARE (cylinder_state = UNINSTALLED). It is loaded from SPIFFS on boots where hub sends SKIP_TARE (cylinder is on the platform).

Observed on this hardware: tare_raw drifts approximately 5–10g per boot due to mechanical creep accumulation. This is expected physics, not a bug. Example observed: -88281.3 → -89234.5 across boots 24–25 (34g drift after platform swap test).

| Situation | Expected behaviour | Action needed |
| --- | --- | --- |
| tare_raw changes by 5–40g between boots | Normal. Mechanical creep, temperature, minor platform settling. | None. System self-corrects. |
| tare_raw changes by 40–200g between boots | Possible platform disturbance or unusual temperature swing. | Monitor. If readings look wrong, fresh tare. |
| tare_raw changes by > 200g between boots | Platform was moved or something changed physically. | Investigate. Fresh tare required. |
| tare_raw same across many boots (SKIP_TARE path) | Normal. Hub correctly preserving tare while cylinder is on. | None. This is correct behaviour. |

# 8. Quick Reference — What to Do In Each Situation

| Situation | Correct action |
| --- | --- |
| Want to add rubber mat / drip tray (before first deployment) | Place it on empty platform. Take tare. Mat is invisible forever. |
| Want to add rubber mat / drip tray (system already running) | Remove cylinder → wait 120s (UNINSTALLED) → wait 60–90s creep recovery → place mat → click Install cylinder (fresh tare) → reinstall cylinder. |
| Want to remove rubber mat / drip tray | Same as above. Remove cylinder first, then mat, then fresh tare. |
| Platform was moved or disturbed | Remove cylinder → wait 120s → wait 60–90s → click Install cylinder (fresh tare) → reinstall cylinder. |
| Swapping entire platform assembly | Remove cylinder → power off node → swap platform → level to 0.5° → power on node → fresh tare fires automatically → re-derive cal_factor if new cells → reinstall cylinder. |
| Just removed cylinder, want to take tare now | Wait 60–90s minimum (10–20 min after long deployment). Let creep recover. Then click Install cylinder. |
| Gas% seems wrong by a large amount | Suspect platform disturbance or wrong tare. Remove cylinder, fresh tare, reinstall. |
| Added something to platform with cylinder still on it | Read wrong. Remove cylinder, remove added item, fresh tare, reinstall. No software fix. |

Gas Cylinder Monitor V1 · Gratian Technologies · Project 13 · Platform Operations Reference · 2026-06-25