# GAS CYLINDER WEIGHT MONITOR

# MASTER DESIGN DOCUMENT (MDD)

## Version 2.0

## PART 3 — MEASUREMENT ARCHITECTURE & CALIBRATION THEORY

# 25. INTRODUCTION

The previous section solved:

“How do we convert weight into gas?”

This section solves:

“How do we obtain trustworthy weight?”

The gas estimation layer is useless if the measurement layer is inaccurate.

Therefore:

Measurement quality is the foundation of the entire system.

# 26. MEASUREMENT PHILOSOPHY

The project adopts a simple principle:

A single reading is evidence.

A stable reading is truth.

## What This Means

The system never trusts:

- One HX711 sample

- One ADC conversion

- One weight measurement

Instead:

Measurements must pass through validation layers before being accepted.

## Design Principle

Nothing is trusted directly from the sensor.

Everything must earn trust.

# 27. COMPLETE MEASUREMENT STACK

The measurement pipeline is:

Raw HX711 Counts ↓ Corrupt-Value Filtering ↓ Clean Samples ↓ Averaging ↓ Stability Detection ↓ Scale Zero Removal ↓ Gross Weight ↓ Steel Removal ↓ Gas Remaining ↓ Percentage

## Importance

Each stage solves a different problem.

Combining stages causes bugs.

Keeping stages separate creates clarity.

# 28. RAW HX711 COUNTS

The HX711 does not output:

- grams

- kilograms

- percentages

It outputs:

24-bit signed ADC counts.

## Meaning

These counts represent amplified bridge voltage.

The values themselves are not meaningful to users.

They are merely sensor observations.

## Example

Raw Count:

835,112

This number has no direct physical meaning.

Calibration converts it into weight.

# 29. CORRUPT VALUE FILTERING

Certain readings are known to be invalid.

Examples:

Communication errors

Timing violations

Sensor faults

Power disturbances

## Rule

Invalid values are rejected immediately.

They never enter averaging.

They never enter storage.

They never influence calibration.

## Principle

Bad data should die early.

Never allow corruption to propagate.

# 30. CLEAN SAMPLE STAGE

After filtering,

the system possesses a stream of valid observations.

These observations are still noisy.

However:

They are no longer obviously invalid.

## Goal

Transform:

Untrusted Samples

into

Candidate Measurements

# 31. NOISE

Noise is unavoidable.

The objective is not to eliminate noise.

The objective is to understand it.

## Sources

Electrical Noise

Mechanical Vibration

Temperature Effects

Contact Resistance

Power Supply Variation

Environmental Disturbance

## Observation

Noise changes.

It is not constant.

A system calibrated today may behave differently tomorrow.

# 32. NOISE CHARACTERISATION

The system measures noise.

It does not guess noise.

## Measurements

Standard Deviation (σ)

Peak-to-Peak Spread

Mean Drift

## Purpose

Noise measurements determine:

Averaging Depth

Detection Thresholds

Health State

Confidence Levels

## Engineering Rule

Noise is a measured quantity.

Not a configuration constant.

# 33. AVERAGING

Averaging reduces random error.

## Principle

Random noise tends to cancel.

Signal tends to persist.

## Result

Averaged readings are more trustworthy than individual samples.

## Important Clarification

Averaging does not remove:

Calibration Error

Drift

Bias

Incorrect Wiring

Mechanical Problems

It only reduces random variation.

# 34. WHY N IS NOT CONSTANT

Many systems choose:

N = 10

or

N = 50

arbitrarily.

This project rejects that approach.

## Correct Approach

Measure Noise ↓ Determine Detection Margin ↓ Derive N

## Benefit

The system adapts to actual conditions.

Not assumptions.

# 35. STABILITY DETECTION

Averaging alone is insufficient.

The object being measured may still be moving.

## Example

User places cylinder.

Weight oscillates.

Scale flexes.

Mechanical settling occurs.

Averaging during motion creates incorrect results.

## Requirement

Measurements must become stable before being trusted.

# 36. STABLE PLATEAUS

The system seeks plateaus.

A plateau is:

A period where weight remains essentially constant.

## Characteristics

Low Variance

Minimal Trend

Sufficient Duration

## Importance

Plateaus represent physical reality.

Transitions do not.

# 37. MEASUREMENT QUALITY STATES

Every measurement receives a quality classification.

GOOD

Stable

Low noise

Valid calibration

DEGRADED

Higher uncertainty

Possible drift

Increased noise

FAILED

Measurement cannot be trusted

## Goal

Consumers of measurements never guess quality.

Quality is explicit.

# 38. THE TWO-TARE MODEL

One of the most important concepts in the project.

## Tare #1

Scale Zero

Represents:

Empty Platform

Purpose:

Remove mechanical/electrical offset

Layer:

Measurement Layer

## Tare #2

Cylinder Steel

Represents:

Empty Cylinder Shell

Purpose:

Separate steel from gas

Layer:

Domain Layer

# 39. WHY THE TWO TARES MUST NEVER MIX

Scale Zero solves:

Sensor offset.

Steel solves:

Gas estimation.

Confusing them creates:

Incorrect percentages

Incorrect gas estimates

Calibration errors

Misleading analytics

## Rule

Every discussion must specify:

Scale Zero

or

Steel Weight

Never simply:

“Tare”

# 40. CALIBRATION

Calibration converts:

Raw Counts

into

Grams

## Purpose

Create a mathematical relationship between:

Sensor Output

and

Physical Reality

Without calibration:

The system only knows counts.

# 41. CAL_FACTOR

The primary calibration parameter is:

cal_factor

Purpose:

Counts → Grams

Important:

cal_factor is hardware-specific.

## Consequence

A calibration derived on:

STM32

cannot automatically be used on:

ESP32-C3

Every platform must establish its own calibration.

# 42. MULTI-POINT CALIBRATION

Single-point calibration is insufficient.

Reason:

Sensors may not behave perfectly linearly.

The preferred method uses:

Multiple Known Weights

across the operating range.

Benefits:

Detect nonlinearity

Improve confidence

Reveal hysteresis

Expose mounting issues

# 43. LINEARITY

Question:

Does one cal_factor remain valid across the range?

Example

100 g

500 g

1000 g

5000 g

10000 g

A perfectly linear system uses the same conversion everywhere.

A non-linear system does not.

# 44. HYSTERESIS

Question:

Does loading equal unloading?

Example

Add 500 g

Measure

Remove 500 g

Measure

If results differ:

Hysteresis exists.

## Consequence

The scale remembers its history.

This must be quantified.

# 45. EXPERIMENT 005

Purpose:

Characterise measurement quality.

Questions Answered

Linearity

Hysteresis

Low-range accuracy

Residual errors

Calibration validity

## Importance

Experiment 005 validates the measurement model itself.

# 46. DRIFT

Drift is the most dangerous long-term error source.

Unlike noise:

Drift does not average away.

Examples

Temperature

Mechanical Creep

Mounting Stress

Component Aging

# 47. THERMAL DRIFT

Temperature changes sensor behavior.

Questions

How much does weight shift per °C?

Does scale zero move?

Does cal_factor move?

These must be measured.

Never assumed.

# 48. LONG-RUN DRIFT

Some errors appear slowly.

Hours.

Days.

Weeks.

Purpose

Understand baseline wander.

Detect creep.

Measure long-term stability.

# 49. DIURNAL DRIFT

Real kitchens experience daily temperature cycles.

Morning

Afternoon

Night

The system must understand:

How the real environment affects measurements.

# 50. SELF-VALIDATION PHILOSOPHY

The project does not trust calibration forever.

Instead:

Measure ↓ Validate ↓ Cross-Check ↓ Trust

Examples

Anchor Events

Known Weights

Long-Term Consistency

Residual Analysis

# 51. SENSOR HEALTH

The measurement layer continuously evaluates itself.

Indicators

Noise Increase

Calibration Drift

Missing Samples

Transport Errors

Unexpected Behavior

## Goal

Detect problems before users notice them.

# 52. ANALOG WIRING DISCIPLINE

The load cell signal exists in the millivolt range.

Small wiring issues become large measurement errors.

Potential Problems

Loose Connections

Jumper Resistance

Mechanical Stress

Poor Contacts

## Engineering Rule

Analog wiring quality directly affects measurement quality.

Treat analog paths as critical infrastructure.

# 53. MEASUREMENT LAYER SUMMARY

The measurement layer exists to answer:

“What is the true weight right now?”

It accomplishes this through:

Filtering

- Characterisation

- Calibration

- Validation

Only after trustworthy weight exists can higher layers perform:

Gas Estimation

Analytics

Prediction

Everything above this layer depends on its integrity.

END OF PART 3