Multi-Cell Load Cell Physics, Wiring, and Electrical Theory

Complete First-Principles Reference

gas-cylinder-monitor | Gratian Technologies | 2026-06-10

Platform: ESP32-C3 SuperMini + GISLAB HX711 + YZC-161A 20 kg load cells

Status: VERIFIED — all equations proven from first principles in this session

# Section 1 — Single Load Cell: Physics from First Principles

## 1.1 What a Load Cell Actually Is

A load cell is not a sensor in the conventional sense. It contains no transistors, no active circuitry, and no digital components. It is a precision metal beam — machined to a specific geometry — with four thin metallic foils bonded to its surface. These foils are called strain gauges.

When weight is placed on the free end of the beam, it bends. The bending is microscopic — on the order of micrometers — but it is enough to stretch or compress the foils. Stretching a metal foil changes its electrical resistance. That resistance change is the entire weight signal.

## 1.2 Stress, Strain, and Hooke's Law

Three equations govern what happens physically when force is applied.

Stress — force distributed over area:

σ = F / A

σ = stress (Pa),  F = force (Newtons),  A = beam cross-section (m²)

Strain — fractional deformation:

ε = ΔL / L

ε = strain (dimensionless),  ΔL = change in length,  L = original length

Hooke's Law — connecting stress, strain, and material stiffness:

σ = E × ε   →   ε = F / (E × A)

E = Young's modulus (~200 GPa for steel)

The critical conclusion: strain is linearly proportional to force. Double the force → double the strain. This linearity is not an approximation — it is fundamental to metal behaviour within the elastic range.

## 1.3 Beam Bending Zones

- Top surface: compressed — fibres squeezed together. Gauges here: R decreases.

- Bottom surface: stretched — fibres pulled apart. Gauges here: R increases.

- Neutral axis (centre plane): zero strain. No gauges placed here.

Key takeaway: A load cell is a precision metal beam. Bending = weight signal. Everything else converts that bending into voltage.

# Section 2 — The Gauge Factor: How Strain Becomes Resistance

## 2.1 The Gauge Factor Equation

A strain gauge is a thin metallic foil bonded to the beam. When it stretches or compresses, its resistance changes. The relationship is:

ΔR / R = GF × ε

GF = Gauge Factor ≈ 2.0 for metal foil gauges

Substituting Hooke's Law (ε = F/(E×A)):

ΔR / R = GF × F / (E × A)          ← THE MASTER EQUATION

This is the master equation. ΔR/R is linearly proportional to force F. Every gram produces a fixed, repeatable resistance change. This is what makes weight measurement possible.

Key takeaway: ΔR/R = GF × F/(E×A). Resistance change is linear in force. This is the foundation of the entire measurement chain.

# Section 3 — The Wheatstone Bridge: Resistance Change to Voltage

## 3.1 The Bridge Circuit

Four gauges are wired in a diamond (Wheatstone bridge) with four nodes:

E+ (Red wire, 3.3V)  — excitation positive, from HX711

E- (Black wire, GND) — excitation negative, from HX711

A+ (signal output)   — read by HX711 Channel A

A- (signal reference)— read by HX711 Channel A

Under load: R1 and R4 increase by ΔR (tension surface). R2 and R3 decrease by ΔR (compression surface).

## 3.2 Deriving Vout — Full Working

The bridge has two voltage divider arms.

Right arm (E+ through R1 to A+, then R3 to E-):

V(A+) = Vex × R3 / (R1 + R3)

= Vex × (R - ΔR) / ((R+ΔR) + (R-ΔR))

= Vex × (R - ΔR) / (2R)

Left arm (E+ through R2 to A-, then R4 to E-):

V(A-) = Vex × R4 / (R2 + R4)

= Vex × (R + ΔR) / (2R)

Differential output:

Vout = V(A+) - V(A-)

= Vex × (R-ΔR)/(2R) - Vex × (R+ΔR)/(2R)

= Vex × [(R-ΔR) - (R+ΔR)] / (2R)

= Vex × (-2ΔR) / (2R)

= -Vex × ΔR/R

Substituting the master equation:

Vout = -Vex × GF × F / (E × A)

PROVEN: Vout is linearly proportional to force F. No assumptions. No approximations.

Key takeaway: Vout = -Vex × GF × F/(E×A). The voltage output of a single load cell bridge is directly proportional to the applied force.

# Section 4 — Thevenin Equivalent of a Single Bridge

## 4.1 Why Thevenin Matters

Thevenin's theorem: any circuit with sources and resistors, seen from two terminals, = one ideal voltage source (Vth) in series with one resistor (Rth). This reduces a complex bridge to two numbers, making multi-cell analysis simple algebra.

## 4.2 Vth — Open Circuit Voltage

Vth = -Vex × ΔR/R = -Vex × GF × F / (E × A)

## 4.3 Rth — Thevenin Resistance

Replace Vex with a short circuit (E+ and E- become same node). Looking from A+ to A-:

Left path:  R2 + R4 = R + R = 2R

Right path: R1 + R3 = R + R = 2R

Two paths in parallel: Rth = 2R ∥ 2R = R

## 4.4 Parallel Resistance Formula — Derived from KCL

Same voltage V across Ra and Rb:

I_total = V/Ra + V/Rb = V × (Ra+Rb)/(Ra×Rb)

I_total = V/R_parallel

1/R_parallel = 1/Ra + 1/Rb

R_parallel = (Ra × Rb) / (Ra + Rb)

Equal case (Ra=Rb=X): R_parallel = X/2

For the bridge: Rth = 2R × 2R / (2R + 2R) = 4R²/4R = R

## 4.5 Complete Thevenin Result

Vth = -Vex × ΔR/R    (open-circuit voltage)

Rth = R              (two-terminal resistance)

Rs  = R/2            (single terminal source impedance)

Key takeaway: Every load cell bridge collapses to Vth + Rth. This is what the rest of the circuit sees.

# Section 5 — Circuit Theory Foundations

## 5.1 What Is a Node

A node is any point where two or more wires connect. It is not a component — it is a connection point. Physical reality: the A+ node is the spot where all White wires (or Green wires for your setup) from all cells are twisted together and connected to HX711 A+. That physical twist IS the node.

## 5.2 Kirchhoff's Current Law (KCL)

Origin: conservation of charge. Charge cannot be created or destroyed. At any node:

ΣI = 0   (sum of all currents at the node = zero)

ΣI_in = ΣI_out

This is a consequence of Maxwell's equations. It is exact, not an approximation.

## 5.3 HX711 Input Impedance

The HX711 differential input is very high impedance. It draws essentially zero current from the A+ and A- nodes. Therefore KCL at the A+ node simplifies to: currents from all cell sources must sum to zero among themselves.

Key takeaway: Node = twist point. KCL = charge conservation. HX711 draws no current. These three facts let us solve for node voltage algebraically.

# Section 6 — Two Cells in Parallel: Full KCL Proof

## 6.1 Physical Wiring — No Junction Box Needed

- Cell 1 Red + Cell 2 Red → twist → HX711 E+

- Cell 1 Black + Cell 2 Black → twist → HX711 E-

- Cell 1 White + Cell 2 White → twist → HX711 A+  (the A+ node)

- Cell 1 Green + Cell 2 Green → twist → HX711 A-  (the A- node)

The twist itself is the junction. No box, no terminal block needed.

## 6.2 KCL at the A+ Node

Each cell = Thevenin source: V1 or V2, with Rs1 = Rs2 = R/2. HX711 draws I=0.

I1 = (V1 - V_node) / (R/2)

I2 = (V2 - V_node) / (R/2)

KCL: I1 + I2 = 0

(V1 - V_node)/(R/2) + (V2 - V_node)/(R/2) = 0

Multiply by R/2:

(V1 - V_node) + (V2 - V_node) = 0

V1 + V2 - 2 × V_node = 0

V_node = (V1 + V2) / 2

PROVEN: node voltage = arithmetic average of the two source voltages. Direct result of KCL + Ohm's Law.

## 6.3 Final Vout for Two Cells in Parallel

V_node(A+) = Vex/2 - Vex×(ΔR1+ΔR2)/(4R)

V_node(A-) = Vex/2 + Vex×(ΔR1+ΔR2)/(4R)

Vout = V_node(A+) - V_node(A-)

= -Vex × (ΔR1 + ΔR2) / (2R)

= -Vex × GF × (F1 + F2) / (2 × E × A)

= -Vex × GF × F_total / (2 × E × A)

PROVEN: Vout is proportional to F_total = F1 + F2. The individual split between cells does not appear in the result.

Key takeaway: KCL at A+ node → V_node = average of sources → Vout proportional to F1+F2. Proven rigorously.

# Section 7 — Why Node Voltage Is an Average, Not a Sum

## 7.1 Physical Intuition

Cell 1 "wants" to push the A+ node to V1. Cell 2 "wants" to push it to V2. They fight each other through equal source resistances Rs1 = Rs2 = R/2. Equal pulling strength → node settles exactly midway. The equal source impedance (from identical cells) makes it a plain average, not a weighted one.

If Rs were unequal (one cell is a different model with different R), the stronger source (lower Rs) would dominate. The result would be a weighted average. But with identical YZC-161A cells, Rs1 = Rs2, so the result is a plain average.

## 7.2 Why Average × Differential = Total

The averaging happens at both A+ and A-. When you subtract:

Vout = V_node(A+) - V_node(A-)

= [Vex/2 - Vex(ΔR1+ΔR2)/(4R)]

- [Vex/2 + Vex(ΔR1+ΔR2)/(4R)]

= -2 × Vex×(ΔR1+ΔR2)/(4R)

= -Vex×(ΔR1+ΔR2)/(2R)

The Vex/2 common terms cancel. The signal terms add. The factor of 2 from the two-terminal differential recovers what the averaging at A+ "lost." Total signal is preserved.

Key takeaway: Average × 2 terminal differential = total. The averaging does not lose signal — it is a balanced circuit that preserves total force information.

# Section 8 — Unequal Load Distribution: Proof the Split Is Irrelevant

## 8.1 Algebraic Proof

ΔR1 = R × GF × F1 / (E × A)    (depends only on F1)

ΔR2 = R × GF × F2 / (E × A)    (depends only on F2)

Vout = -Vex × (ΔR1 + ΔR2) / (2R)

= -Vex × [GF×F1/(E×A) + GF×F2/(E×A)]

= -Vex × GF × (F1 + F2) / (2 × E × A)

= -Vex × GF × F_total / (2 × E × A)

F1 and F2 appear only as their sum F_total. The split has no effect on Vout.

## 8.2 Numerical Verification

| Split | F1 (g) | F2 (g) | ΔR1+ΔR2 (relative) | Vout |
| --- | --- | --- | --- | --- |
| 50/50 | 500 | 500 | 1000×k | -Vex×1000k/(2R) |
| 80/20 | 800 | 200 | 1000×k | -Vex×1000k/(2R) |
| 99/1 | 990 | 10 | 1000×k | -Vex×1000k/(2R) |
| 100/0 | 1000 | 0 | 1000×k | -Vex×1000k/(2R) |

All four rows produce identical Vout. The split is algebraically invisible.

Key takeaway: Vout = -Vex×GF×F_total/(2EA). Any distribution of F_total gives the same reading.

# Section 9 — N Cells in Parallel: General Formula

## 9.1 KCL Extended to N Equal Sources

Σ (Vi - V_node)/(R/2) = 0   for i = 1 to N

ΣVi = N × V_node

V_node = (V1 + V2 + ... + VN) / N

Vout = -Vex × GF × F_total / (N × E × A)

## 9.2 Effect on CAL_FACTOR

| Cells (N) | Sensitivity relative to 1 | Approx CAL_FACTOR (raw/g) |
| --- | --- | --- |
| 1 | 1× | ~105 |
| 2 | 1/2× | ~52 |
| 3 | 1/3× | ~35 |
| 4 | 1/4× | ~26 |

These are approximate values for the YZC-161A on ESP32-C3 platform. Actual values must be measured by calibration on the specific assembled hardware.

Key takeaway: N cells: Vout proportional to F_total always. Sensitivity reduces by N. CAL_FACTOR must be re-derived for each N.

# Section 10 — Why Parallel, Not Series

## Problem 1 — Broken Excitation Reference

In parallel: both E+ pins connect to same Vex. Common ground reference for both bridges. In series: Cell 2's E- floats relative to Cell 1. No stable reference. Requires separate Vex per cell.

## Problem 2 — Offset and Noise Stacking

In parallel: each cell's offset is common-mode on both A+ and A-. Cancelled by HX711 differential measurement. In series: offsets of all cells add. Temperature drift in one cell corrupts the reading of all cells.

## Problem 3 — Common-Mode Voltage Violation

In parallel: both A+ and A- held near Vex/2 (~1.65V for 3.3V). Within HX711 input spec. In series: stacking outputs pushes absolute voltage higher — potentially outside HX711 input spec.

Key takeaway: Series breaks excitation sharing, stacks offsets, and may violate HX711 input spec. Parallel solves all three. Parallel is the only correct topology.

# Section 11 — CAL_FACTOR and Wiring Topology

## 11.1 The CAL_FACTOR Equation

grams = (raw_reading - tare_raw) / CAL_FACTOR

CAL_FACTOR = net_raw_counts / known_weight_grams

CAL_FACTOR is not a universal constant. It encodes the sensitivity of the specific wiring topology.

## 11.2 Topology Mismatch Rule

reported = actual × (N_cal / N_meas)

N_cal = number of cells connected during calibration

N_meas = number of cells connected during measurement

| N_cal | N_meas | Reported weight | Error |
| --- | --- | --- | --- |
| 4 | 4 | Correct | PASS |
| 1 | 4 | 25% of actual | FAIL — 75% low |
| 4 | 1 | 400% of actual | FAIL — 300% high |
| 3 | 4 | 75% of actual | FAIL — 25% low |
| 2 | 4 | 50% of actual | FAIL — 50% low |

ALWAYS calibrate with all cells connected in their final configuration. If topology changes, re-derive CAL_FACTOR from scratch.

Key takeaway: CAL_FACTOR encodes the topology. Topology changes → re-calibrate. N_cal must equal N_meas.

# Section 12 — Three Cells vs Four Cells

## 12.1 Electrical Argument (Both Work)

3 cells in parallel sums correctly: Vout proportional to F1+F2+F3. 4 cells in parallel sums correctly: Vout proportional to F1+F2+F3+F4. Electrically, any N works.

## 12.2 Mechanical Argument (Shared Platform Only)

For a shared platform: stability requires the cylinder's center of mass (CoM) to stay inside the support polygon (shape formed by connecting support points). If CoM exits the polygon, the platform tips. A lifted cell reads zero. The electrical sum becomes wrong.

| Configuration | Support polygon | Platform coverage | Unstable zones |
| --- | --- | --- | --- |
| 1 cell | Point | ~0% | All of platform |
| 2 cells | Line | ~0% (tips perpendicular) | Half the platform |
| 3 cells (triangle on square) | Triangle | ~58% | Corner zones |
| 4 cells (corners of square) | Full rectangle | 100% | None |

For a production household product with a shared platform: 4 corner cells cover 100% of the platform. No tipping is geometrically possible under any cylinder placement.

Key takeaway: Electrically 3 and 4 cells both sum correctly. Mechanically, 4 corner cells on a square platform guarantee 100% stable zone coverage.

# Section 13 — Independent Plates vs Shared Platform

## Case 1 — Independent Plates (Your Experiment)

Each cell has its own plate. Weights placed independently. No mechanical coupling.

Example: Cell1=180g, Cell2=340g, Cell3=500g, Cell4=empty

HX711 reads: 180+340+500+0 = 1020g

Serial monitor shows: 1020g. Not the individual weights. Always the sum.

## Case 2 — Shared Platform

One rigid plate on all cells. One object placed on platform. Platform distributes weight across cells.

Example: 1020g cylinder, split 400/350/270 across 3 cells

HX711 reads: 400+350+270 = 1020g

Serial monitor shows: 1020g. Same as Case 1 for the same total.

## Are the Two Cases Different?

NO. In both cases HX711 reads the sum of all forces. Same total = same reading. The distribution is invisible.

What HX711 cannot tell you with a single chip: which cell has what weight, where on the platform the object is, whether load is even or uneven. One HX711 = one number = total force only.

Key takeaway: Independent plates or shared platform — same total = same reading. HX711 is a total-force sensor, not a distribution sensor.

# Section 14 — Experiment Design: Sum Invariance Verification

## Your Weights

Phone = 234g    Speaker = 229g    Adapter = 92g    Wood = 34g

Total = 589g

## Critical Precondition

All 4 cells must be wired and connected during calibration AND all 8 measurement configurations. Disconnecting any cell after calibration changes N_meas → N_cal/N_meas mismatch → systematic wrong readings.

## Test Matrix

| Config | Cell 1 | Cell 2 | Cell 3 | Cell 4 | Expected | Tests |
| --- | --- | --- | --- | --- | --- | --- |
| C1 | Phone 234g | Spkr 229g | Adpt 92g | Wood 34g | 589g | Sum invariance |
| C2 | Wood 34g | Adpt 92g | Spkr 229g | Phone 234g | 589g | Sum invariance |
| C3 | ALL 589g | empty | empty | empty | 589g | 100/0/0/0 extreme |
| C4 | empty | ALL 589g | empty | empty | 589g | 0/100/0/0 extreme |
| C5 | Ph+Sp 463g | empty | Ad+Wo 126g | empty | 589g | Two-pair split |
| C6 | Phone 234g | empty | empty | empty | 234g | Partial loading |
| C7 | Ph 234g | Sp 229g | empty | empty | 463g | Partial loading |
| C8 | empty | empty | empty | empty | 0g | Tare drift |

Pass criterion: each configuration average within ±5g of expected. C1 through C5 must all read within ±5g of each other.

# Section 15 — Locked Rules and Handoff Notes

## HANDOFF NOTE: CAL_FACTOR Recalibration

CAL_FACTOR must be re-derived whenever topology changes:

- Adding or removing any load cell → re-derive

- Moving from N-cell test to M-cell production → re-derive

- Replacing any cell (even same model) → re-derive

- Reassembling platform with different geometry → re-derive tare AND CAL_FACTOR

The single-cell values (106.7 raw/g on STM32, ~113 raw/g on ESP32-C3) are VOID on the 4-cell platform. Expected 4-cell value: ~26–28 raw/g. Must be measured, never computed.

## Wiring Rule

All Red wires   → twist → HX711 E+

All Black wires → twist → HX711 E-

All White wires → twist → HX711 A+

All Green wires → twist → HX711 A-

Twist = junction. No junction box needed. Verify wire colors against your cell datasheet — convention varies by manufacturer.

## Corrupt Filter Rule — All Three Required

LONG_MIN  (-2147483648) : timeout or sign-extension artifact

-1        (0xFFFFFF)    : read mid-conversion (all bits HIGH)

0x7FFFFF  (8388607)    : positive saturation (wrong pin/broken cell)

Never omit any of the three. Each catches a different failure mode.

## Float Rule

Use float throughout on ESP32-C3. Do not use double. The double sum=0 bug was confirmed on STM32U585 — status on ESP32-C3 unverified. Float is safe and sufficient.

## Timing Rules

- delayMicroseconds(1) after every SCK edge — both HIGH and LOW, data bits AND gain pulse

- delay(110) minimum between reads — HX711 at 10Hz needs 100ms per conversion

- noInterrupts() wrapping full 24+1 pulse sequence — FreeRTOS must not interrupt mid-read

- millis() pacing guard at TOP of loop() — never inside a state case

- No while() loops inside state machine cases — one sample per loop() iteration

# Quick Reference — All Key Equations

PHYSICS:

σ = F/A                              (stress from force)

ε = ΔL/L                             (strain definition)

ε = F/(E×A)                          (Hooke's Law)

ΔR/R = GF × F/(E×A)                 (master equation)

SINGLE CELL BRIDGE:

V(A+) = Vex × (R-ΔR)/(2R)           (right arm divider)

V(A-) = Vex × (R+ΔR)/(2R)           (left arm divider)

Vout  = -Vex × ΔR/R                 (differential output)

Vout  = -Vex × GF × F/(E×A)         (output vs force)

THEVENIN:

Vth = -Vex × ΔR/R                   (open-circuit voltage)

Rth = R                              (two-terminal resistance)

Rs  = R/2                            (single terminal impedance)

PARALLEL RESISTANCE:

R_p = (Ra × Rb) / (Ra + Rb)         (two resistors)

R_p = X/2  (when Ra = Rb = X)       (equal resistors)

KCL NODE (N equal sources):

V_node = (V1+V2+...+VN) / N         (arithmetic average)

N CELLS IN PARALLEL:

Vout = -Vex × GF × F_total/(N×E×A)  (any distribution)

CALIBRATION TOPOLOGY:

reported = actual × (N_cal / N_meas)

WEIGHT CONVERSION:

grams = (raw - tare_raw) / CAL_FACTOR

gas-cylinder-monitor | Gratian Technologies | 2026-06-10

All equations proven from first principles in Claude.ai learning session

Upload to: ~/ArduinoApps/gas-cylinder-monitor/docs/ and project knowledge base